## Slide 1: SchwingNet: Naming a Vertebral Level From One C-arm Exposure, Without Counting From C2
Gregory J. Schwing, MD, PhD
Department of General Surgery, Detroit Medical Center and Wayne State University, Detroit, MI
CNS 2026 Annual Meeting · Washington, DC · [session and date pending]

## Slide 2: Where this is going
A real lateral C-arm exposure; each body named from its shape, with its probability.
1
Surgeons count vertebrae on the C-arm. On a transitional spine the count is wrong
2
Today's tools copy a planning label onto the film. The planning read is wrong on the same spines
3
SchwingNet reads the level from the shape of the bones in one exposure
4
Tested on a C-arm it never saw: how often it is right, and when it stops
5
What is missing: enough transitional patients to prove it, and the CT census that will find them

## Slide 3: A level is counted, not marked, and the count is wrong often
Strong et al. 2022, Figure 3: intraoperative fluoroscopy, needle at the T10 pedicle, counted from the sacrum. J Neurosurg Spine 36:792 (PMC10193475); permission pending.
Half of spine surgeons report a wrong-level operation in their career (Mody 2008; Groff 2013)
Registry of 34,544 operations: 0.12 %, thoracic 0.42 % (Furuya 2026)
The first exposed level is wrong in 4–15 % before the check film (Ammerman 2006; Dablouk 2019)
No neurophysiological check between T1 and L2; verification protocols miss a third (Longo 2012)
The remedy on offer is another count.
1. Mody MG, et al. Spine 2008;33:194. 2. Groff MW, et al. World Neurosurg 2013;79:585. 3. Furuya M, et al. Bone Joint J 2026;108-B:673. 4. Dablouk MO, et al. Br J Neurosurg 2019. 5. Ammerman JM, et al. Surg Neurol 2006. 6. Kirshblum SC, et al. J Spinal Cord Med 2011. 7. Longo UG, et al. J Bone Joint Surg Br 2012;94-B:1546.

## Slide 4: Transitional anatomy changes the count
Labelled CT, public de-identified records (CTSpinoPelvic1K). Bar, 20 mm.
8.8 % of 3,147 full-length radiographs lack 24 presacral segments (Sabri 2025)
The count was wrong in 38 of 71 patients with a transitional vertebra (Nakagawa 2017)
The truth is on CT. No public radiograph or C-arm frame carries a human-read T13 or Castellvi grade
8. Sabri SA, et al. Neurosurg Focus 2025;58(6):E13. 9. Nakagawa T, et al. 2017. 10. Castellvi AE, et al. Spine 1984;9:493. Figure: SchwingNet manuscript (in preparation), Figure 1.

## Slide 5: What a wrong level costs, and what avoiding it costs
Bettis et al. 2026, Figure 1: National Practitioner Data Bank malpractice payment reports for surgical never events, 1991–2025 (Patient Saf Surg, open access; permission pending).
17 % of events lead to legal action; one in five Spine Section surgeons has been sued (Mody 2008; Groff 2013)
81 % of litigated cases paid; median never-event payment $75,220 (Goodkin 2004; Bettis 2026)
Avoiding it: a gold fiducial placed the day before, 52 min and 183 mGy per patient, 179 patients at one centre (AJNR 2023)
Derived, unpublished: roughly 420–1,900 US wrong-level operations a year; indemnity alone about $8–37 million.
11. Goodkin R, Laska LL. Surg Neurol 2004;61:323. 12. Bettis T, et al. Patient Saf Surg 2026. 13. Machin JT, et al. Eur Spine J 2018;27:2693. 14. AJNR Am J Neuroradiol 2023;44:618. 15. Childers CP, Maggard-Gibbons M. JAMA Surg 2018;153:e176233.

## Slide 6: The label on the film is only as right as the planning read
Top: De Silva et al. 2016, Figure 1 (Spine; PMC5045321). Bottom: Manbachi et al. 2018, Figure 6, labels on a 13th rib and a lumbarised S1 (Ann Biomed Eng; PMC9107073). Permission pending.
LevelCheck registers preoperative CT labels onto the film: 100 % placement, 97.6 % surgeon reads (De Silva 2016; Manbachi 2018)
Scored against the imported labels. On a transitional spine the label is only as right as the planning read
Registration-free naming on real C-arm images: 73 % (Zhang 2023); cadaver long-length images 88 % (Huang 2022)
Nothing published names a level from one exposure alone, and nothing is tested on a transitional spine.
16. Otake Y, et al. Phys Med Biol 2012;57:5485. 17. Lo SFL, et al. Spine 2015. 18. De Silva T, et al. Spine 2016. 19. Manbachi A, et al. Ann Biomed Eng 2018;46:1548. 20. Peh WCG, et al. 1999. 21. Zhang R, et al. J Xray Sci Technol 2023. 22. Huang Y, et al. Comput Methods Programs Biomed 2022.

## Slide 7: The planning read fails where the count fails
A lumbarization case with a true L6. Left, ground truth; right, the segmenter's output.
An open-source whole-body segmenter, unmodified, on 1,152 public CTs (Wasserthal 2023)
Normal anatomy: L5 Dice 0.94. Lumbarization: every label slides one level, L5 Dice 0.29
L6 emitted: 0 of 1,152. Confident, well formed, one level off, no error signal
Not All Spines Are Created Equal, NeurIPS 2026
23. Wasserthal J, et al. Radiol Artif Intell 2023;5:e230024. 24. Schwing GJ, et al. Not All Spines Are Created Equal. NeurIPS 2026 Datasets and Benchmarks (accepted; retrospective). 25. CTSpinoPelvic1K, doi.org/10.5281/zenodo.22139642.

## Slide 8: Hypothesis and purpose
Four phenotypes at the lumbar borders, aligned on the sacrum (CTSpinoPelvic1K, public, de-identified).
Hypothesis: each transitional vertebra has a local shape a network reading a few bodies can learn
Purpose: build it, test it on an unseen C-arm, and count the transitional persons a test needs. Targets set before scoring: 90 % naming, 85 % variant sensitivity, 97 % detection
Investigational. Decision support beside the surgeon's count, never a replacement for the check film.

## Slide 9: What exists to learn from: every source, one map
Supplement Figure S1a: t-SNE of ImageNet features; renders sit apart from every real source. Two kV-imager sources relabelled generically.
11,638 images from 18 public sources, embedded by a general image network and laid out in two dimensions
Films · C-arm fluoroscopy · kV imagers · CT renders, plain · CT renders styled to a device
4,093 CT renders, 9,108 films with levels, 4,276 real exposures; 194 transitional persons
The intraoperative C-arm set (148 cases, cone-beam CT truth, 2023) is in no training set
Retrospective study of public collections, analysed April–September 2026; sources, licences and download locations per release in the manuscript. VinDr-SpineXR under a PhysioNet data use agreement.

## Slide 10: What SchwingNet does: inputs and outputs
1 · Find
bodies and sacrum boxed;
no level name
2 · Name
each body named from its shape,
probability beside it
3 · Check
against the count from the sacrum:
agree, a name; disagree, a stop
In: one exposure. Out: a name with its probability for every body, the count from the sacrum, a flag when they disagree
No end of the spine needed; no label imported from another study
Method withheld pending university invention review; weights and code with the paper

## Slide 11: How it was tested
Real lateral frames, calls by the cross-validated model; green right, red wrong, orange unscored (labellers disagree on the body). Needle tip on the named L4.
148 intraoperative cases, public, acquired November–December 2023; truth from each patient's cone-beam CT
Three readings: never trained on this C-arm; fine-tuned by case-grouped folds; by acquisition-day folds
Exact level per detected body; intervals clustered by case; every number read from a result file by the build

## Slide 12: A sixth lumbar vertebra has a shape of its own
The lumbosacral window of the embedding; SchwingNet manuscript, Figure 2b.
13,009 validation crops, 1,612 people: the nearest crop of another person is the same level in 74.9 %
L6 separates from L5 (silhouette 0.35) as far as L1 from T12 (0.32)
75 of 98 L6 crops named L6
Preliminary: ten L6 persons, on the split that chose the checkpoint

## Slide 13: One lateral exposure of a C-arm the network never saw
C-ArmLSpine, public intraoperative exposures, 2023; SchwingNet manuscript, Figure 4a–b.
Zero-shot, lateral: 595 of 617 bodies exact (96.4 %; 92.3, 98.4)
Anteroposterior zero-shot 43.7 %; fine-tuned on other cases 96.4 %
Three bodies in view 93.1 %; one body 74.8 %: a name where there is nothing to count

## Slide 14: Two readings of one frame, and a stop where they disagree
Same frame. (c) Ordinals from the detected sacrum, named in the ordinary five-lumbar alphabet. (d) A frame one CT labeller reads six-lumbar: shape and count disagree on body 1; the read stops.
Count from the sacrum 97.0 %, shape alone 97.7 %, together 98.2 % (1,009 bodies)
Stopped on 23 of 233 frames: five of the six errors and all four L6 calls
A disagreement is reported, not hidden
How the readings are combined and when the read stops are withheld pending the invention review.

## Slide 15: Where the frames sat mattered more than how many
SchwingNet manuscript, Figure 5c–d; render sweep on 40 unseen subjects.
Unseen device: 80.5 % of bodies found; 93.9 % on truth boxes inside the field
Its own frames: a plateau from 24 (c). One view buys that view only
Renders at two poses find 53 % between them; the same count spread over the arc, 77 % (d)

## Slide 16: How many people it takes, and which view they were shot in
Exact level against training persons; SchwingNet manuscript, Figure 5b.
Film naming rose with training persons: NHANES II 78 → 93 %, BUU-LSPINE 65 → 97 %, CSXA 94 → 99.8 %
C-arm lateral: 93.7 % with no C-arm person. Anteroposterior: 33.8 % with none, 90.8 % with five
Training both views cost neither (98.1 and 97.0 %)
● NHANES II
● BUU-LSPINE
▲ CSXA
● C-arm, lateral
■ C-arm, anteroposterior
○ lateral-trained, read on anteroposterior
✕ both views, all persons

## Slide 17: Transitional naming on a real exposure: the one case
C-ArmLSpine case 42, public; SchwingNet manuscript, Figure 4d, with the probabilities the system printed on each box.
A patient one CT labeller reads six-lumbar, the other five. Lateral frame, cross-validation fold
Shape names bodies 5 to 2 as L2–L5 (0.89–1.00); the count says L1–L4. Body 1: shape none (0.02), count L5
The call is L6 for the column; body 1 is left unnamed and the read stops
Three such patients exist, six frames; scored in cross-validation only. A description, not a test
No accuracy on this slide is a wrong-level surgery rate. Evidence base absent: transitional naming on real exposures is unverified.

## Slide 18: What the census would let me claim
n: the transitional persons a screen of every collection in the census should find, from published prevalence; a case counts only where a film exists. Today there are 7 candidates on real exposures, none verified.
Each cell: the lowest accuracy the data would still support at 95 % confidence if the network named those n persons with that share of errors (Wilson bound). 157 error-free persons are the fewest that clear 97.6 %.
97.6 % is what surgeons scored reading 62 difficult films with the LevelCheck overlay (Manbachi 2018). With 870 L6 persons the network may miss 12 and still sit above it; with all 2,069 transitional persons, 36. The 322 L6 of the screen running now allow 2.
Beating it outright needs the paired comparison on the same persons, a separate study; this slide is the size of the claim the census makes possible.

## Slide 19: Limitations and next steps
Limitations
One device; its truth is two automatic labellers agreeing
148 cases may not be 148 people; intervals may be too narrow
Real frames sit at two poses; one patient's rotational acquisition is the only real arc
Retrospective, public data; nothing prospective; not a cleared device
Next steps
Verify the transitional candidates; run the 35-person, zero-error test
A second C-arm, with obliques, held out; phantom, cadaver, then human series
Bodies found on one patient's real rotational acquisition, 301 frames:
● detector that also saw this device's spot frames
■ ▲ other devices' renders and frames, this one held out
◆ the reported detector, this one held out
Supplement Figure S5: frames at four angles (bar, 50 mm) and detection against view angle at confidence 0.25.

## Slide 20: Conclusions
Exact level on the unseen C-arm.
The count fails on transitional anatomy; registered labels and CT segmenters inherit the miscount
SchwingNet: 96.4 % zero-shot on the lateral view, 98.2 % coupled, and a stop on 23 of 233 frames holding five of six errors
Transitional naming on real exposures: unverified. 35 persons needed; the census is under way
A name from shape, checked against the count, with a stop where they disagree
Evidence status: investigational; retrospective, public de-identified data; no prospective evaluation; not a cleared device.

## Slide 21: Acknowledgments and data availability
With thanks to Patrick Schwing (LSU Health Sciences Center New Orleans); Miraziz Ismoilov and Nizar Alnabahneh (Radiology) and Loren Schwiebert (Computer Science), Wayne State University; and the annotators Ashley Schehr, Annika Tekumulla, Margret Khoushi, Ryan Christian, Dane Hubers, Faris Mahjoub, Hassan Saad, Mia Sooch, Sathyagopal Siddapureddy, Michael McLellan and Jerick Kim
Data: CTSpinoPelvic1K (doi.org/10.5281/zenodo.22139642) and the public C-arm, radiograph and CT collections under their own licences; all de-identified
Code and weights: github.com/OpenSpineConsortium
Funding: NIH NIGMS F30GM147909. No commercial support.
Questions
gregory.schwing@med.wayne.edu
openspineconsortium.com